| 題名: | PSPICE Implementation for Designing the Acoustical Matching Layers of Ultrasonic Piezoelectric Transducer |
| --- | --- |
| 作者: | Yeong-Chin Chen;Long Wu;Yuan-Tsong Chang;Cheng-Liang Huang |
| 關鍵詞: | PSPICE;ultrasonic diagnostic;matching layer;wide-band;high frequency;high transmitting power;piezoelectric transducer;KLM model;lossy acoustical transmission line |
| 日期: | 1999 |
| 上傳時間: | 2010-04-15T05:41:50Z |
| 摘要: | Ultrasound has become an effective and practical tool for obtaining medical information from within the human body without exposing the patient to X-rays. The emphasis of this paper is on the design of low-loss (high transmitting power) and wide-band transducers used for ultrasonic diagnostic. The approach we have used consists of impedance matching the front face of the transducer to the propagating medium and air-backing the rear face. By tuning the thickness and acoustical impedance of the matching layer, transducers with wide-band and high-transmitting-power can be obtained. PSPICE code of the KLM model is implemented to predict precisely the performance of the matched transducers. Since the transducer's characteristics of wide-band width and high-transmitting-power are somewhat in conflict in the design criteria, a matching-layer thickness that is a little bit shorter or longer than one quarter wavelength is proposed as a good trade-off between the two criteria. Good agreement between the simulation results and experimental results has been achieved. |